DATA ACQUISITION SYSTEM DATALOG 20/90/140

Acquisition, monitoring and data processing

- INPUT and OUTPUT

- 1 to 2000 channels

- High accuracy system

- Statistics and calculations

- PCMCIA

- MODBUS/RTU

The DATALOG series is intended to measure, handle, process, monitor and record analogue and digital signals from all common types of physical sensors. The systems can be supplied in various configurations with a choice of input and output boards as well as options such as built-in display and printer.

Whatever system chosen, and thanks to the measurement metrological quality, the programming power offered by the internal functions, the DATALOG systems are well adapted to numerous applications, such as:

- Stand-alone data acquisition system without PC.

- Sensor calibration system.

- Testing bench.

- Conditioning chamber and autoclave validation.

- PC input/output interface.

- Acquisition module for programmable controller or standard supervisor.

Functions

Embedded software..................

The DATALOG series provides the user with a large amount of programming power.

This can be used or not depending upon the work to be carried out.

Channels........................................

Each module can address 2 000 channels.

These can be real channels (input or output) or virtual channels performing computations.

Monitoring................................

4 programmable thresholds per channel.

A specific relay output or, if required, a conditional processing may be associated to each threshold.

Mathematical operations.......

Any mathematical, Boolean and statistical computations may be defined on or between channels. 100 linearization table with 40 couples of counts each may be accessed by the user (measured value associated with computed value), thus allowing corrections of sensors.

Conditional processing......…….....

IF, THEN, DO are accessible and authorise changes to constants, thresholds, tasks, i.e. condition scanning.

Channel scanning..........……........

The system is organised by “tasks” (up to 100) which define the inputs/outputs scanning and result output to peripherals.

Storage and PCMCIA.......……….......

The DATALOG are equipped (as standard) with 6 storage memories of 8 000 measurements (1 per task over the first six tasks). All the DATALOG with the optional keypad are equipped with a PCMCIA slot authorizing loading of various configurations and large capacity storage.

Configuration and processing …...

All the DATALOG can be configured using PC software which enables the operator to collect the data stored for PC processing (curves, delayed computations, export to spreadsheets).

Software............................................

In addition to the configuration and processing software, a real time supervisory software is available for networking.

Various utilities (automatic collecting of data via modem) a DLL library and a LABVIEW driver complete the range.

Configuration

(1)Full, half- or quarter- bridge gauges. Need a power supply board which occupies 1 I/O slot.

(2) RS 485 or RS 232C network interface. MODBUS/RTU (binary) protocol. Up to 32 modules can be linked. D < 1 200 m.

Functions

Analogue input boards AN 5885-AN 5906-AN 5905-ATC 017

These boards are for universal inputs.

Each channel is configurable depending on the quantity to be measured. The connection is performed over a removable screw connector for 0.5 to 2.5 mm2 wires.

AN 5885.......................................

10 input channels authorising the following measurements:

DC voltages and currents.

AC voltages and currents.

Resistance using 3- or 4- wire configuration.

Platinum and Nickel RTDs.

Thermocouples

Dry contacts

Strain gauges (Need the AN 3700 board).

AN 5906.......................................

10 input channels. Same functions as for the AN 5885, except Strain gauges measurements and voltage is limited to 60V.

ATC 017......................................

This board is to be used for energy source measurements. Resistors mounted in series on the inputs ensure the protection against accidental switching short-circuits.

It has 10 input channels for measuring as follows:

DC voltages and currents, AC voltages and currents.

AN 5905.......................................

20 input channels intended for measuring as follows:

DC voltages and currents, AC voltages and currents, Thermocouples, Dry contact

Switching..............................

It is performed over the AN 5885, AN 5900 and ATC 017 boards by using 3-wire dry contact relay.

Differential resistance ≤ 40 mΩ.

Stray emf: ± 2.5 mV.

Life ³ 108 operations.

For the AN 5905, the switching is static and performed by means of optomos and the voltage between channels is limited to 60 VDC or AC.

For the other boards, the voltage between channels is limited to 150 VDC or AC.

Scanning speeds

Three integration times are programmable channel per channel. This criterion acts on the conversion and settling time, the lowest speed gives the maximum resolution and accuracy.

(1) RU = Representation Unit which cor responds to the minimum deviation between two displayed values. E.g.: for a decimal display changing 3 counts per 3 counts, RU = 3.

(2) Analogue-to-digital converter, 12 bits with successive approximations. This speed is not available for strain gauges.

For lower speeds, converter ± 690 000 counts.

DC Voltage

AN 5906 has same ucertainties as AN 5885, but a volatge range upto 60 V.

At 20 meas/s, the uncertainty is slightly reduced with regard to 7 meas/s.

At 50 meas/s, the uncertainty over one year is ranging 4 RU (measurement resolution).

At 100 meas/s, the uncertainty over one year is ranging 12 RU (measurement resolution).

The “process current” ranges need an external 50 Ω - 0.1% shunt.

Rejection level at 50 Hz ± 1% over 60 mV range at 7 and 20 meas/s:

- Normal mode > 60 dB

- Common mode > 140 dB.

Input impedance:

100 MΩ over ranges < 6 V

10 MΩ for upper ranges.

Temperature coefficient:

- From 0 to 35°C:(0.0001% + 0.5 RU)/°C.

- From 35 to 50°C: (0.001% + 1 RU)°C.

Repeatability between two channels: <Constant C.

AC Voltage

T-RMS value measurement (AC + DC mode). The accuracies (uncertainties) are given for a sine wave voltage without DC component.

Measurement minimum scanning time: 2 seconds.

Measurement range: 4 to 110% of range from 40 to 400 Hz.

Crest factor influence (CF): 1% for CF = 3.

Repeatability between two channels: < C.

Temperature coefficient:(0.05% + 0.5 RU)°C.

1) The uncertainty is given in ± (% rdg + C) at 23°C ± 5 °C with rdg = reading and C = Constant.

Input impedance 100MΩ on ranges<6V 10MΩ fro upper ranges

Nota: boards AN 5905 and AN 5906 limited to 60V

Resistance

Not available for AN 5905

Connection 3 or 4 wires

(1)Inn ± (%R + C) à 23 ± 5°C. with R = Reading et C = constant.

The uncertainties at 20 meas/s are the same as for 7 meas/s.

At 50 meas/s, the uncertainty over one year is ranging 4 RU.

At 100 meas/s, the uncertainty over one year is ranging 12 RU.

When measuring with 3- wire configuration, add 100 mΩ to the uncertainties above and, if required, the possible line unbalance.

Measurement current: 1mA for ranges from 60 Ω to 3 kΩ and 10μA for 30kΩ and 300 kΩ ranges.

Permissible line resistance: < 100Ω per wire.

Repeatability between two channels:

3-wire < (100 mΩ + 3 RU)

4-wire < (10 mΩ + 3 RU).

1

RTD's

Unavailable with the AN 5905 input board.

Connection: 3- or 4- balanced wire.

RTD table according to IEC 751.

(1)Inn ± (%R + C) à 23 ± 5°C. with R = Reading et C = constant.

The uncertainties at 20 meas/s are the same as for 7 meas/s.

At 50 meas/s, the uncertainty over one year is ranging 2 RU.

At 100 meas/s, the uncertainty over one year is ranging 3 RU.

When measuring with 3- wire configuration, add 0.25°C to the uncertainties above and the line unbalance value converted in °C.

Measurement current: 1 mA.

Permissible line resistance ≤ 100Ω per wire.

Temperature coefficient:(0.002% + 0.0025°C)/°C.

Repeatability between two channels: ≤ 0,05°C with 4 wires, ≤ 0,2°C with 3 wires.

Thermocouples

Three reference junction compensation modes are programmable: without RJC, with built-in RJC (1 Pt 100 per board), with remote RJC (1Pt 100 per board).

Thermocouple linearization according to IEC Publication 584-1.

Expressed in ± (% rdg + C) at 23°C ± 5°C with rdg = r eading and C = Constant..

The uncertainties above are given for an RJC at 0°C.

Using the built-in RJC, add to the uncertainties above:

0.2°C for the AN 5885 et AN5906 boards

0,3°C for AN 5905 board.

Repro Repeatability between two channels:

- over a same board: 1 RU, between two different boards: 1 RU plus the RJC error.

Input resistance ³100M Ω

Permissible line resistance: KW per wire

Strain gauges AN3700+AN5885

Sensor power......................................

A specific power board common to all “gauge” channels occupies one standard slot in the system.

Available voltage: 2 V or 10 V switched to each bridge measured.

Measurement current: 1 mA or 8 mA switched to each half- or quarter- bridge measured.

Voltage and current are selected by switches located on the gauge power board. Three compensation gauges can be connected to three quarter- bridge configurations.

Measured signals applied to the standard inputs of the analogue input boards as follows: one channel per quarter- bridge, two consecutive channels per half- or full bridge. Free channels can still be used for other purposes.

Sensors: minimum 120Ω gauges or transducers, maximum voltage drop permissible in one gauge: 3.2V.

Permissible unbalanced voltages compatible with the 30 mV and 300mV ranges of the system voltmeter.

Measurement rate: 7 meas/s, 20meas/s or 50 meas/s.

Note: The system can measure in full bridge configuration with externally supplied power, in this case, the gauge power board is not necessary.

Digital input board AN5886

This board counts pulses or measures frequencies of ten periodic signals.

Each channel is coil-insulated from other channels of the data acquisition system.

Input voltage for any channel:

Level 1: CH1 > + 2.9 V.

Level 0: CH0 < + 1 V.

Maximum usable voltage: ± 50 VDC or AC peak.

Input current for level 1: 1.6mA typical.

Input insulation: 100 VDC or AC RMS

Frequency inputs ............................

Uncertainty: 0.02% + 2 RU.

Temperature coefficient: 0.0004%/°C.

NB: No filter is provided on frequency inputs. A cyclic ratio of approx. 50% (square signal) is advisable to obtain

Stable measurements but the pulse duration should always be above 5μs. max. between each channel.

Counter inputs .....................................

Count positive transitions on the inputs.

User selectable bounce filtering time:

500μs and 1 ms (typical) by a switch located on the digital input board. The filter on channels 1 and 2 can be eliminated to permit high-speed counting.

A jumper on the digital input board allows the use of channel input 10 as a trigger to validate or inhibit other counter inputs on the same board.

Analogue output board AN 5888

The board is made up of 5 digital/analogue converters (12 bits) each providing two output quantities, a voltage - 10 to + 10 V and a current 0-20 mA or 4- 20 mA.

As voltage and current quantities come

Voltage output............................

Minimum permissible load resistance: 5 kΩ.

Current output............................

- Internal power: Maximum permissible load resistance: 500 Ω.

Ten Open circuit voltage: 12V. from the same converter, they cannot be used simultaneously.

Voltage and current channels of the same board are coil-insulated and also are insulated from other system input/outputs.

- External power:

Supply voltage: 15V ≤V supply.≤ 50V.

Max. permissible load resistance Rmax = (V supply - 2)/0.02.

Max. permissible common mode voltage between analogue channels and other system I/O channels: 150 VDC or VAC peak Average processing time: 20 ms.

Signal risetime to resistance load: 20μs/V.

Maximum load capacity: 10μF.

Temperature coefficient: 0.01% per °C.

Dry relay output board AN 5887

Capacity: 10 bistable relays per board.

Wired to removable screw terminal connector.

Contact resistance load: 48 V or 1 A or 30 Ω.

Number of operations: 5.105 in the following conditions.

Each contact is either common, on or off.

The relays are controlled according to system program, for example:

- on a specific alarm limit action;

- remote controlled from one of the communication interfaces

- on a conditional processing.

. Communication Interfaces

The DATALOG are equipped in standard with two communication interfaces, each configurable in RS 232 by means of a DB9 female connector or RS 485 using a 5-pin screw terminal block.

Data format: 8 bits, 1 stop bit, no parity.

Addresses are programmable from 1 to 15 and transmission rates from 9 600 to 56 000 bauds.

P1 interface ........................................

MODBUS RTU protocol either as master (for reading other instruments) or as slave.

P2 interface ........................................

Operation configurable in ASCII RS232 (printer management for example) or MODBUS RTU (slave only).

Internal Printer

Option only available with DATALOG 90 and 140.

Thermal printer, 24 characters per line.

Printing speed: 1.5 line per second.

Paper roll: 80 mm wide, 50 m long, and diameter 65 mm.

PCMCIA Interface

The DATALOG with keypad version, are equipped with a PCMCIA interface, type 2. This interface enables the operator to read and write on memory cards of RAM STATIC technology and FLASH card of ATA type. These cards allow to save and load configuration files as well as to create result files.

Main Characteristics

Operating conditions ........................

Reference range: 23°C ± 5°C.

Operating nominal range: 0°C to 50°C, 20% to 80% RH noncondensing.

Power requirements .........................

100 to 230 VAC ± 10%, 50 Hz to 60 Hz.

Consumption: 40 VA.

Input available for 11 V to 28 VDC supply.

Optional battery pack with built-in charger.

Life: 3 hours approximately.

Unregulated output available for sensor supply. Rated voltage 24 V, current 100mA.

Presentation........................................

The systems are delivered with a carrying handle to be removed for panel mounting with optional accessories. A protective cover located at back of the system enables the operator to measure voltages above 60 V in accordance with the EMC and low voltage directives.

The systems can be delivered with various options (refer to the table).

Weight................................................

From 3 kg to 9 kg depending on models and options..

Dimensions...................................

Software

Various utilities for Windows from 3.1 to NT are available.

Logidat ..........................................

Used to configure the DATALOG and process the values stored in memory from compatible PC. Easy to use thanks to menu-guided operation.

Processing the result files offers possible delayed computations, graphic form, sort and export to spreadsheets.

Connection to the systems can be performed by modems.

Visulog ......................................

Real time Monitoring software for DATALOG; ot shows measurement in realtime as mimic diagrams, values or curves.

Storage of values and alarms in traceability files and in Log files.

Data management and exploitation can be performed without stopping acquisition.

Refer to the specific documentation of Visulog for more details.

Email sending on alarm.

.

Utilities...................................

In order to simplify all specific developments, a wide range of utilities is available: Visulog DDE links with Windows environment.

DLL library, 16 and 32 bits.

Labview driver.

Instructions pour commander

System 2 I/O without keypad and display                                D2AO

System 2 I/O with keypad and display                                   D2CO

System 9 I/O without keypad and display                                D9AO

System 9 I/O without keypad and display but with rechargeable battery        D9AB

System 9 I/O with keypad and display (1)                                D9CO

System 9 I/O with keypad, display and rechargeable battery (1)               D9CB

System 9 I/O with keypad, display and internal printer (1)                   D9CI

System 14 I/O without keypad and display                               D14AO

System 14 I/O without keypad and display but with rechargeable battery       D14AB

System 14 I/O with keypad and display (1)                               D14CO

System 14 I/O with keypad, display and rechargeable battery (1)              D14CB

System 14 I/O with keypad, display and internal printer (1)                  D14CI

10-channel analogue input board                                       AN 5885

10-channel simplified analogue input board                              AN 5906

10-channel protected analogue input board                              ATC 017

20-channel analogue input board, 2-wire                                 AN 5905

Strain gauge board                                                  AN 3700

10-channel digital input board                                         AN 5886

10-channel relay output board                                         AN 5887

5-channel analogue output board                                       AN 5888

Accessoires

Removable terminal block for I/O boards, 10 channels                     ER 48276-000

Removable terminal block for I/O boards, 20 channels                     ER 48402-000

PCMCIA flash card                                                ATC013

Shunt for process current measurement                                 ER 44007-024

Set of 10 paper rolls for DATALOG                                   ATC030

IEEE mechanical adapter for ATC 006                                 ATC020

Voltmeter and CPU for DATALOG                                   ATC031

Supply for DATALOG                                             ATC032

Protection cover for DATALOG 20                                   ATC023

Protection cover for DATALOG 90                                   ATC024

Protection cover for DATALOG 140                                  ATC025

Rack kit for DATALOG 20                                          ATC026

Rack kit for DATALOG 90                                          ATC027

Rack kit for DATALOG 140                                         ATC028

12-thermocouple cone for autoclave                                   ATC033

12-platinum probe cone for autoclave                                  ATC034

Specific cone for autoclave                                          ATC035

Software

Windows configuration software                                     LOGIDAT

Real-time monitoring windows software                               VISULOG

Additional licence for VISULOG                                     LICVISU1