DIGITAL PANEL METERS ICO 62

Electrical measurements of process temperature measurements 10 000 points

The function of measurement, the range, the thresholds as well as the specific scalings are programmable. The configured parameters are protected from all inopportune modifications.

These aircraft are equipped with two relays of thresholds.

- Format 96 x 24 mm

- Function, ranges, thresholds and

- scale programmable.

- 2 thresholds with exit relay.

Functions

Measure...............................................

The ICO 62 allow measurements of voltage, running, resistance of process and temperature measurement using resistor probes Pt or of thermocouples.

They are equipped:

• blocking of measurement

• memorizing of the maximum and minimum.

Scaling...............................................

The value to be displayed is entered using the keyboard, while the value of corresponding measurement can either be introduced in the same way or measured by the instrument.This possibility is interesting for the corrections of errors on temperature ranges.

Thresholds...............................................

Programming of two thresholds and their hysteresis.

Capacity breaking of the associated relay 250 V - 1 A or 60 VA to 50 H z, 150 V - 1 A or 3 0 W in D.C.

In the course of measurement, each threshold can be visualized and modified.

Temperature measurements...............................................

(1)For a type of determined sensor, the commutation of range and the change of resolution can be automatic.

Possibility of going beyond of range. Display in C or F.

(2) ± (% reading +°C numbers)on 1 year without junction of reference.The typical error brough by the junction of reference AN8002 is of 0,3°C.

Electrical measurements of process...............................................

(3) Extended from measurement : - 40 to + 110 % range

(4) ± (% dreading +a number of units)on 1 year.

Protection of the programming...............................................

Two possible levels by software:

- For all the programming,

- Protection or not of the threshold values.

General characteristics

Display...............................................

From -1999 to 9999 with indication, LED red,yellow or green,of 14 mm.

Speed of measurement : 2,5 measures/s.

Temperature coefficient...............................................

≤ 10 % accuracy/°C.

Climatic Environnement...............................................

Nominal field : 0 to 50°C, 20 to 80 % HR without condensation.

Acceptable maximal voltage...............................................

600 V~ on the ranges 5 V and 50 V;

250 V~ on the ranges 50 mV, 500 Mv and thermocouples;

20 V~ on the entries of measurement of resistance

Common mode...............................................

Rerjection ³ 110 dB to 50 Hz on the range 50 mV.

Acceptable maximal voltage : 250 V ~.

Rejection of series mode...............................................

³ 60 dB to 50 Hz on the range 50 mV.

Isolated power supply...............................................

230 V~ (± 10 %) 50 to 60 Hz.

Others power supply:

5 V- ± 10 %,

9 to 18 V-,

18 to 36 V-,

36 to 72 V-,

24 V~ ± 10 %,

48 V~, ± 10 %,

115 V~± 10 %.

Presentation...............................................

Case ABS. Weight : approximatly 300g.

Sealing : IP 40.

Connecting...............................................

Removable screwing connectors supplied with instrument.

Dimension

Accessories

External shunt...............................................

AN 8005 : 10 W, 500 mV, accuracy 0,1 %. Skewerable module on the connector of entry for measurements in 4-20 my (current max.: 100 my).

Dimensions : 25 x 27,5 x 14,8 mm.

Compensation of junction of reference...............................................

AN 8002 : modulate screwed on the back, ransportable connector block remotely.

Typical error : 0,3°C.

Max.error : 0,5°C.

Dimensions : 25 x 27,5 x 14,8 mm.

Protective circuit...............................................

Circuit RC of protection of the contacts of alarm for cut on inductive circuits (250 V~ max.) ER42062.

Instructions to order