High accuracy scanner 高精度扫描仪SHP 101

SHP101台式尖端仪器能够扫描直到12个温度传感器。当使用精密温度计PHP601时，热电偶和RTDs能被测量没有信号完整性或在性能方面的损失。每个区域或许被配置或者作为2个测量通道（热电偶或电压测量）或者作为一个4线测量（RTDs）。

- 计量扫描仪

- 热电偶和RTDs。

- RS 232

- 杰出的测量同质性

- 完全与PHP601兼容。

描述..........................................

- 编程和处理：通过使用12键的键盘、RS232链接或可选的IEEE488链接。

- 传感器的连接用螺丝接线端接受4mm插头（碲铜接线端，以非常低的热电偶效应）。

- 阻抗转换，用3或4线配置。

- 热电偶转换，有或没有参考连接点补偿。

- 输出可用于直接连接到高精度温度计PHP601.

应用..........................................

- 扫描热电偶，TRDs和低电平的电信号。

- 温度传感器的校准。

- 映射。

- 使用作为伏特计或欧姆计输入扩展。

功能

The SHP 101 operates as follows:

SHP101操作如下：

- 利用前键盘手工操作。

- 遥控控制或者通过一个PHP601（直接RS232命令）或通过一个控制器（RS232或IEEE接口）

命令..........................................

通道和2或4线配置结构可被选择。

配置结构..........................................

- 6个RTD测量通道，4线配置结构或。

- 12伏电压（热电偶）测量通道，2线配置结构或。

- N 4线通道(0 ≤ N ≤ 6)和2 x (6 – N) 2 线通道。

连接..........................................

输入：4mm碲铜接线端。

电压和参考输出：碲铜接线端。

RTD和参考连接点输出：5针连接器。

开关..........................................

继电器，以非常低的电动势和接触电阻。

杂散电动势：0,2 μV 最大值。

接触电阻：300 mΩ 最大值。

在一个通道的两线之间接触电阻差异最大值：40 mΩ。

在输入区域之间的偏离：± 0.02°C 最大值。

分断能力：60 V, 100 mA, 1 W 最大值。

参考点..........................................

不确定性由于内部的参考点：± 0.05°C 最大值。

温度系数：± 0.005°C/°C 最大值。

一般性指标说明

运行条件..........................................

遵守IEC发行的359：运行类别I。

参考量程：23 ± 1°C。

相对湿度：45 to 75%。

运行的标称量程：0 到 50°C, RH 20 to 75% 不凝固的。

运行量程限值：- 10 to + 55°C, RH 10 to 80% 不凝固的。

存储和运输限值量程：- 30 to + 60°C。

介绍..........................................

ABS台式类型壳提供一个倾斜保证/手柄。

尺寸: 225 x 110 x 310 mm.

重量: 2 kg.

电源需要: 115 V or 230 V ± 10 %, 50 Hz to 400 Hz, 20 VA max.

版本

基础版本包括市电电源和RS232连接。

附件

- Cable, 2-meter length with LEMO socket at one end and RTD free connection at the other end (ACL 4603).

- LEMO socket to be wired (ER 48379).

- PHP-SHP link cable with LEMO sockets at both ends. One is for the PHP/SHP external reference junction compensation, the other is for measuring resistance over the scanner (ACL 4601).

- Connection leads, 1-meter length, PHPSHP equipped with 4-mm dia. plug; very low emf cable for connecting thermocouple or voltage type signals.

订购说明

Scanner                       SHP 101-1扫描仪

Scanner with IEEE 488 interface   SHP 101-3扫描仪带接口

Accessories附件..........................................

Soft carrying case                           AN 6901 软的携带式仪器箱

RS 232 cable, 9-pin male/female               AN 5875 电缆

IEEE 488 cable, 2-meter length                AN 5836 电缆2米长

LEMO socket to be wired                     ER 48379 接有电线的插座

PHP/SHP connection cable                    ACL 4601 连接电缆

Cable, 2-meter length with LEMO and free ends   ACL 4603 电缆2米长

Standard sensor with plugs                    AN 5847B 标准传感器带插头

Reference sensor with plug                    AN 5848-2000A 参考传感器带插头

Standard thermocouple                       T 1200 标准热电偶

Low emf 2-wire cord                         PEM 40317 低电动势2线卡